LXLC(R)-250~300(mm)

Removable element woltman cold (hot) water meter

Application:

--Measuring the total volume of cold(hot) water passing through the pipeline.

-NOTE:This type of water meter can be used for a reading transmission system as equipped with pluse.

Feature:

-- Dry-dial, Magnetic drive

----Vacuum sealed register ensures the dial keep free from fog and Keep the reading clear in a long term service;

--Removable element unit and strong universality, easy installation and maintenance,

--Measuing high flow rate with low head loss.

-- Remote transmission device can be added upon request,such as Reed Swith or Hall, Pulse Output: DN250-300mm  10000L/P、1000L/P

--Technical data conform to international standard ISO 4064..

Working condition

Water temperature: T30 T50 T90

Water pressure class: MAP16  MAP10

Pressure-loss Class: △p63

Maximum permissible errors

(1) From minimum flow-rate(Q1) inclusive to transitional flow-rate(Q2) exclusive±5%.

(2) From transitional flow-rate(Q2) to overload flow-rate(Q4):±2%;

(Hot water meter ±3%)

### MAIN TECHNICAL SPECIFICATIONS

| Type | Size | Class | Q4 Overload Flow | Q3 Permanent Flow | Q2 Transitional Flow | Q1 Min. Flow | Min. Reading | Max. Reading |
| --- | --- | --- | --- | --- | --- | --- | --- | --- |
|  |  |  | m3/h |  |  |  | m3 |  |
| LXLC-250 | 10" | R50 (R80) | 787.5 | 630 | 20.16 | 12.6 | 0.02 | 9999,999 |
| LXLC-300 | 12" |  | 1250 | 1000 | 32 | 20 | 0.02 | 9,999,999 |

### DIMENSIONS

| Type | Size |  | L Length |  | H Height | Connecting flange |  |  |
| --- | --- | --- | --- | --- | --- | --- | --- | --- |
|  |  |  |  |  |  | φD1 Outside diameter | φD2 Bolt circle diameter | Connecting Bolts (n-M) |
|  | mm |  |  |  |  |  |  |  |
| LXLC-250 | 10" | 450 |  | 434 |  | 395 | 350 | 12-M20(1.0MPa) |
|  |  |  |  |  |  | 405 | 355 | 12-M24(1.6MPa) |
| LXLC-300 | 12" | 500 |  | 480 |  | 445 | 400 | 12-M20(1.0MPa) |
|  |  |  |  |  |  | 460 | 410 | 12-M24(1.6MPa) |

NOTE: The flange dimension conforms to ISO7005-2:1988 standard. Order for products of special requirements is also accepted.

| LXLC-250～300 REMOVABLE ELEMENT WOLTMAN WATER METER |  |  |  |
| --- | --- | --- | --- |
| Code | Description | Material | QTY |
| 1 | Screw For Sealing | Stainless Steel | 1 |
| 2 | Cover With Lid | Assembly | 1 |
| 3 | Sealed Register | Assembly | 1 |
| 4 | Screw M16×45 | Stainless Steel | 1 |
| 5 | Fixed plate | PPO | 1 |
| 6 | Adjusting Nut | Brass and Silicon Rubber | 2 |
| 7 | Flange Cover | Cast Iron | 1 |
| 8 | O-ring | Silicon Rubber | 1 |
| 9 | Bush | Brass | 1 |
| 10 | Gasket | Synthetic Rubber | 1 |
| 11 | Adjusting Lever | Brass | 1 |
| 12 | Connecting Lever | Nylon | 1 |
| 13 | Adjusting Plate | Nylon | 1 |
| 14 | Support and gear | Nylon and pom | 1 |
| 15 | Turbine Component | Assembly | 1 |
| 16 | Rectifier | Nylon | 1 |
| 17 | Screw | Brass | 1 |
| 18 | O-ring | Synthetic Rubber | 1 |
| 19 | Body | Cast Iron | 1 |
| 20 | Screw For Sealing | Stainless Steel | 2 |
| 21 | Screw M16X45 | Stainless Steel | 7 |
| 22 | Gasket | Stainless Steel | 8 |
| 23 | Screw for fixing | Stainless Steel | 4 |
| 24 | rings | Stainless Steel | 2 |
| 25 | Gasket | Cast Iron | 1 |
| 26 | Upper Bearing Plate | Brass | 1 |
| 27 | Upper Bearing | Nylon With Graphite | 1 |
| 28 | Magnet Component | Component | 1 |
| 29 | Sheath | Nylon | 2 |
| 30 | Transmission Shaft | Stainless Steel | 1 |
| 31 | Support | Nylon | 1 |
| 32 | Turbine Shaft | Tungsten Steel | 1 |
| 33 | Screw | Brass | 1 |
| 34 | Nut | Brass | 1 |

----------------------------------------------------------------------------------------------------------------------

LIANYUNGANG HYTTEN METER CO.,LTD.

Address:  No.9 Road XinLong Economic and Technological Development Zone Xinpu

Lianyungang, Jiangsu, China.

Tel:      ++86-518-81069628   81069629

Fax:     ++86-518-81069630             Postal code:   222000

E-mail:982801898@qq.com;hytten@hyttenmeter.com

Web:    http://www.hyttenmeter.com //www.hyttenwatermeter.com