Operating Instruction of Ultrasonic water meter

◆ Summary

The product is an intelligent measuring device, can realize accurate measurement for flow rate. It has good appearance,easy installation, accurate measurement,stable operation,strong ability on anti-pollution and anti-corrosion,long service life, less pressure loss, economical and practical characteristics, suit for meterage on life water and central A/C system water etc.

◆ The principle and main usage

- The principles and components

This measuring device are mainly composed of temperature sensor, flow sensor, and a calculator.Temperature sensor measuring the temperature of the water,flow sensor measuring volume of water flowing through a pipe by the time difference,the two data are collected to integral calculation and calculate the used water volume, store and display.

- The main Usage.

Mainly used for measuring the flow rate of water , such as life water, central A/C water system etc.

| DN |  | DN20 | DN25 | DN32 | DN40 |
| --- | --- | --- | --- | --- | --- |
| Dimension | L | 130 | 160 | 180 | 200 |
|  | L1（With connection pipe） | 235 | 281 | 305 | 330 |
|  | W | 125 | 125 | 125 | 125 |
|  | H | 86 | 89 | 100 | 120 |
| Connection Thread |  | R3/4 | R1 | R1 1/4 | R1 1/2 |
| Table of thread |  | G1 | G1 1/4 | G1 1/2 | G2 |

◆  Dimension

DN（20～40）

| DN | Length mm | Width mm | Connecting flange GB/T17241.6-2008 |  |  |
| --- | --- | --- | --- | --- | --- |
|  |  |  | Flange Diameter D1 | Threaded hole Central Diameter D2 | Connecting thread n--M |
| DN50 | 200 | 135 | 165 | 125 | 4—M16 |
| DN65 | 200 | 150 | 185 | 145 | 4—M16 |
| DN80 | 225 | 185 | 200 | 160 | 8—M16 |
| DN100 | 250 | 205 | 220 | 180 | 8—M16 |
| DN125 | 250 | 245 | 250 | 210 | 8—M16 |
| DN150 | 300 | 270 | 290 | 240 | 8—M20 |
| DN200 | 350 | 320 | 340 | 295 | 8—M22 |

DN（50～200)

◆  Installation Instruction

Small DN

Big DN

◆ Normal Display

| Display No. |  | Data Name | Display pattern |  |
| --- | --- | --- | --- | --- |
| Ordinary State |  | Dormant State | No Display |  |
|  | 1 | Integrated Flux | X.X m³ |  |
|  | 2 | Instantaneous flow velocity | XX.XX m³/h |  |
|  | 3 | Accumulated Working time | XXXXX h |  |
|  | 4 | Temperature | XX.XX ℃ |  |
|  | 5 | Table Address | XXXXXXXX |  |
|  | 6 | Current Date | XX XX XX |  |
|  | 7 | Current Voltage | X. XX V |  |
|  | 8 | Version | E XXXXX.X |  |
|  | Under the Cumulative flow display,Keeping encourage the touch the centralization for 8 seconds and into the historical memory, display as follows: |  |  |  |
|  | 1 | The cumulative amount of hot and cold |  | X KW·h/GJ |
|  | 2 | Integrated Flux |  | X．X m3 |
|  | Keep touch encourage centralization for 8 seconds again withdrew from the historical data storage condition |  |  |  |
|  | Under instantaneous flow velocity,Keep Touch the centralization for 8 seconds into the verification status. |  |  |  |
|  | 1 | Flow Inspection |  |  |
|  | 2 | Rate of Flow Inspection |  |  |
|  | Keep touch encourage centralization for 8 seconds again withdrew the verification status |  |  |  |
| No action after enter a dormant state for 5 Minutes , without any display。 |  |  |  |  |

◆ Main technical parameter

| DN（mm） | 20 | 25 | 32 | 40 | 50 | 65 | 80 | 100 | 125 | 150 | 200 |
| --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- |
| Max.Flow Qs(m3/h) | 3.12 | 5 | 7.87 | 12.5 | 31.25 | 50 | 78.75 | 125 | 200 | 312.5 | 500 |
| Permanent Flow Qp(m³/h) | 2.5 | 4 | 6.3 | 10 | 25 | 40 | 63 | 100 | 160 | 250 | 400 |
| Transitional Flow Qt(m3/h) | 0.03 | 0.05 | 0.08 | 0.13 | 0.32 | 0.51 | 0.8 | 1.28 | 2 | 3.2 | 5.12 |
| Min.Flow Qi(m3/h) | 0.02 | 0.03 | 0.05 | 0.08 | 0.2 | 0.32 | 0.5 | 0.8 | 1.28 | 2 | 3.2 |
| Max.Reading (m3) | XXXXXX．XX |  |  |  |  |  |  |  |  |  |  |
| Accuracy Class | 2 |  |  |  |  |  |  |  |  |  |  |
| Pressure loss | ≦0.063MPa |  |  |  |  |  |  |  |  |  |  |
| Max.working pressure | 1.0MPa |  |  |  |  |  |  |  |  |  |  |
| Temperature Grade | T30 |  |  |  |  |  |  |  |  |  |  |
| Temperature Resolution | 0.01℃ |  |  |  |  |  |  |  |  |  |  |
| Environment class | B(Indoor) |  |  |  |  |  |  |  |  |  |  |
| Battery life | >5+1 years |  |  |  |  |  |  |  |  |  |  |
| Installation | Horizontal or Vertical |  |  |  |  |  |  |  |  |  |  |
| Measuring Media | Cold water |  |  |  |  |  |  |  |  |  |  |
| Temperature Sensor | Pt1000 |  |  |  |  |  |  |  |  |  |  |
| Display | 8 Digital LED |  |  |  |  |  |  |  |  |  |  |

◆ Requirement on Installation

The installation location of water meter all parts(including flow sensor components, temperature sensor parts and calculator parts) should be avoid exposure in the sunlight,flooding,freezing,chemical and electromagnetic pollution and convenient for remove and meter reading.

- The Installation of Water meter

- Should be clean thoroughly system pipeline and clear away flax.sandstone and other sundry before installation, to avoid causing meter failures.

- Horizontal or Vertical installation

- The direction of flow should be in same direction with arrow marked on the body.

- The pipe of diameter of two ends to match the body diameter.

- There is Straight pipe that should be 10 times of DN before body and 8 times of DN behind body(DN:Nominal Diameter).

- Making a seal between connection nut and pipe, flow sensor and platinum resistor after finish installation.

- The installation of temperature sensor.

- The temperature sensor should be installed in hole of body testing temperature and seal.

- The lead of temperature sensor can not increase/reduce/change casually.

◆  Daily Maintenance

- Should be check regularly running status and battery voltage is normal or not after finish installation and using.

- When the meter running time reach 6 years or less than 6 years, but display shows data is not clear or voltage display "V", namely the battery should be replaced by the specialized technical personnel, so as not to affect the normal work of the heat meter.

- Touch the temperature interface for 8 seconds into query mode, if the fourth screen show”0000014d” that means the pipe is not full(empty pipe).

◆  Guarantee Clause

Free repair within one year from delivery date and maintenance during life-long, but the following situation will not warranty.

- The parts/seal/logo was be opened and destroyed.

- The parts were damaged by human.

- The parts were sunlight exposure,flooding,freezing,chemical pollution.

- Not clean the pipe or too much impurity in pipe before installation, lead to damage of flow sensor.

- The lead of temperature sensor is broken, or display excessive rotate to lead to the lead of flow sensor is broken.

- The fault and damage caused that not choose suitable product model.

Note: The battery of Ultrasonic water meter should be replaced by special person, User can not do it by himself, should contact with manufacturer directly if undervoltage occur.

◆ Transportation and storage.

- Transportation: Should not be affected directly by rain,frost,fog during transportation.

Place it conform to “UP” mark, and avoid extrusion/collision damage. The transportation of  the environmental conditions according to standard JB/T9329.

- Storage: the mat of product at least 30 cm above the ground, should not be less than 1m from wall and should not be less than 2m from heating equipment. Environment temperature 5 ℃ ~ 55 ℃,relative humidity is less than 80%.Should be no acid, alkali, inflammable, explosive, toxic chemicals and other corrosive gases and goods in warehouse,should prevent intense electromagnetic fields and direct sunlight.

Lianyungang Hytten Meter Co.,Ltd.

Address:  No.8 Hanguo Development Zone,Luoyang, Dunshang Town,Ganyu District

Lianyungang, Jiangsu, China.

Tel:      ++86-518-85907891

Fax:     ++86-518-85907892             Postal code:   222000

E-mail:  hytten@hyttenmeter.com

Web:    http://www.hyttenmeter.com //www.hyttenwatermeter.com